## Slide 1
Notes: 1

## Slide 2
Atomic Force Microscopy and Scanning Probe Microscopy
1981 - Invention of the STM
1986 - Invention of the AFM
Binnig G, Quate CF, Gerber C. Atomic force microscope. Phys. Rev. Lett. 1986 Mar 3;56(9):930-933
Notes: 2

## Slide 3
Atomic Force Microscopy and Scanning Probe Microscopy
Probe
Surface
Notes: 3

## Slide 4
Atomic Force Microscopy and Scanning Probe Microscopy
Probe
Surface
Probe
Surface
Electrons
STM (Scanning tunneling microscope)
Notes: 4

## Slide 5
Atomic Force Microscopy and Scanning Probe Microscopy
Probe
Surface
Surface
Puranik et al. Plant Cell, 2014
STM (Scanning tunneling microscope)
AFM (Atomic Force Microscope)
Force
Notes: 5

## Slide 6
Techniques in Structural Biology
Dimensions in life science – what technique is appropriate
S. Scheuring, AFMBioMed Summer School, Marcoule 2011
Notes: 6

## Slide 7
Techniques in Structural Biology
Dimensions in life science – what technique is appropriate
S. Scheuring, AFMBioMed Summer School, Marcoule 2011
LENGHTSCALE
TIMESCALE
Notes: 7

## Slide 8
Outline
1) Introduction to Scanning Probe Microscopy
2) The Force
3) Methods
Cantilevers
Static and Dynamic AFM
4) Instruments
5) Advanced and novel AFM methods
Notes: 8

## Slide 9
Interaction Forces at nanoscale
Force
Tip-sample distance
Force
Repulsive
Repulsion due to overlap of
tip-sample
electronic shells
Notes: 9

## Slide 10
Interaction Forces at nanoscale
Force
Tip-sample distance
Force
Attraction due to
Van der Waals, capillarity….
Attractive
Notes: 10

## Slide 11
Interaction Forces at nanoscale
Richard P.
Feynman
(1918 – 1988)
“… If, in some cataclysm, all of scientific knowledge were to be destroyed, and only one sentence passed on to the next generations of creatures, what statement would contain the most information in the fewest words? I believe it is the atomic hypothesis (or the atomic fact, or whatever you wish to call it) that all things are made of atoms—little particles that move around in perpetual motion, attracting each other when they are a little distance apart, but repelling upon being squeezed into one another...”
Notes: 11

## Slide 12
Interaction Forces at nanoscale
Tip sample Interaction Force depends on
Tip and sample geometry
Tip and sample nature (chemical affinity, charges)
Tip and/or Sample stiffness
Environment (liquid, air, gas, vacuum)
Notes: 12

## Slide 13
Interaction Forces at nanoscale
Tip and/or Sample stiffness
Tip-sample distance
Force
Tip-sample distance
Force
RIGID
SOFT
SOFT
Notes: 13

## Slide 14
Interaction Forces at nanoscale
Tip and/or Sample stiffness
Tip-sample distance
Force
Tip-sample distance
Force
RIGID
SOFT
SOFT
Daniel J Müller, Dimitrios Fotiadis, Simon Scheuring, Shirley A Müller & Andreas Engel Electrostatically balanced subnanometer imaging of biological specimens by atomic force microscopy Biophys J, 1999, 76 (2): 1101-1111
IONS ARE
IMPORTANT!!!
Environment: liquid + salt
Notes: 14

## Slide 15
Outline
1) Introduction to Scanning Probe Microscopy
2) The Force
3) Methods
Cantilevers
Static and Dynamic AFM
4) Instruments
5) Advanced and novel AFM methods
Notes: 15

## Slide 16
Cantilevers
=
Mass m
Force
Force
Cantilever
Stiffness k
Sample
Sample
Notes: 16

## Slide 17
Cantilevers
Tip-sample distance
Force
Attractive
Mass m
Mass m
Notes: 17

## Slide 18
Cantilevers
Mass m
Mass m
Tip-sample distance
Force
Repulsive
Notes: 18

## Slide 19
Cantilevers
Mass m
Force
Top view
Side view
Stiffness k
Notes: 19

## Slide 20
Cantilevers
Stiffness k: from 0.006 N/m to 100 N/m
Extrinsic property
Materials: Silicon and Silicon Nitride
Length L: from 10 µm to 500 µm
k must be
calibrated !!
E = Young’s modulus [Pa] intrinsic property
Notes: 20

## Slide 21
Cantilevers and harmonic oscillators
Mass m [Kg]
Finteraction [N]
Stiffness k [N/m]
Damping δ [N s /m]: environment, viscosity
Z
m
Notes: 21

## Slide 22
Cantilevers and harmonic oscillators
Mass m [Kg]
Finteraction [N]
Stiffness k [N/m]
Damping δ [N s /m]: environment, viscosity
Z
m
Notes: 22

## Slide 23
IMAGE ACQUISITION AND FORCE SPECTROSCOPY
Scan
direction
Imaging
Notes: 23

## Slide 24
IMAGE ACQUISITION AND FORCE SPECTROSCOPY
Scan
direction
Imaging
Force
Spectroscopy
Separation
Notes: 24

## Slide 25
Static AFM mode (contact mode)
STATIC AFM: tip in permanent contact with the sample
Phonograph
Notes: 25

## Slide 26
Static AFM mode (contact mode)
STATIC AFM: tip in permanent contact with the sample
Phonograph
Notes: 26

## Slide 27
Static AFM mode (contact mode)
STATIC AFM: tip in permanent contact with the sample
Phonograph
Notes: 27

## Slide 28
Static AFM mode (contact mode)
STATIC AFM: tip in permanent contact with the sample
Phonograph
Cantilever
deflection
Notes: 28

## Slide 29
Static AFM mode (contact mode) – The closed loop
PIEZO: X/Y/Z scanner
Cantilever deflection
Z scanner position
Scan direction
Notes: 29

## Slide 30
Static AFM mode (contact mode) – The closed loop
PIEZO: X/Y/Z scanner
Cantilever deflection
Z scanner position
Notes: 30

## Slide 31
Static AFM mode (contact mode) – The closed loop
PIEZO: X/Y/Z scanner
Cantilever deflection
Z scanner position
Notes: 31

## Slide 32
Static AFM mode (contact mode) – The closed loop
PIEZO: X/Y/Z scanner
Cantilever deflection
Z scanner position
Notes: 32

## Slide 33
Static AFM mode (contact mode) – The closed loop
PIEZO: X/Y/Z scanner
Cantilever deflection
Z scanner position
Notes: 33

## Slide 34
Static AFM mode (contact mode) – The closed loop
PIEZO: X/Y/Z scanner
Cantilever deflection
Z scanner position
Constant
deflection
Notes: 34

## Slide 35
Static AFM mode (contact mode) – Operational scheme
Cantilever
Sample
X/Y/Z Scanner
Optical
system
Cantilever deflection
Magic box
z = 0, z = 0, = 0  = kΔz
Hook
Law
Notes: 35

## Slide 36
Static AFM mode (contact mode) – Force curves
Adapted from Polymer Science, Chapter 4, U. Maver et al., Intech, 2013
Δz
Notes: 36

## Slide 37
Adapted from Polymer Science, Chapter 4, U. Maver et al., Intech, 2013
Δz
H.-J. Butt, B. Cappella and M. Kappl,
Surface Science Reports, 59, 1-152, 2005
Static AFM mode (contact mode) – Force curves
Notes: 37

## Slide 38
Static AFM mode (contact mode) – Force curves
Mechanics
Costa et al., PLOSONE, 2014
Notes: 38

## Slide 39
Static AFM mode (contact mode) – Force curves
Mechanics
Costa et al., PLOSONE, 2014
Molecular recognition
Ligand
Receptor
Notes: 39

## Slide 40
Static AFM mode (contact mode) – Force curves
Mechanics
Costa et al., PLOSONE, 2014
Molecular recognition
Ligand
Receptor
Protein unfolding
F. Rico et al.
Science,
342(6159),
741-743, 2013
Luigi
Bubacco
website
Notes: 40

## Slide 41
Outline
1) Introduction to Scanning Probe Microscopy
2) The Force
3) Methods
Cantilevers
Static and Dynamic AFM
4) Instruments
5) Advanced and novel AFM methods
Notes: 41

## Slide 42
Dynamic AFM mode (Amplitude Modulation mode) – AM-AFM
Mass m [Kg]
Finteraction [N]
Stiffness k [N/m]
Damping δ [N s /m]: environment, viscosity
Z
m
Notes: 42

## Slide 43
Dynamic AFM mode (Amplitude Modulation mode) – AM-AFM
Mass m [Kg]
Finteraction [N]
Stiffness k [N/m]
Damping δ [N s /m]: environment, viscosity
Z
Resonance frequency
m
Notes: 43

## Slide 44
= =
= F / (-mω2 + iδω + k)
Dynamic AFM mode – Harmonic Oscillator
Notes: 44

## Slide 45
= F / (-mω2 + iδω + k)
z = a + ib
Modulus: |z| =
Phase: φ = 2arctan()
Dynamic AFM mode – Harmonic Oscillator
Notes: 45

## Slide 46
= F / (-mω2 + iδω + k)
z = a + ib
Modulus: |z| =
Phase: φ = 2arctan()
A = |z| =
Phase = arctan(φ) =
Dynamic AFM mode – Harmonic Oscillator
A = Amplitude
Frequency
Notes: 46

## Slide 47
Dynamic AFM mode – Harmonic Oscillator (Q factor)
Amplitude
Frequency
Δ
Q factor
High Q
Notes: 47

## Slide 48
Dynamic AFM mode – Harmonic Oscillator (phase)
F
A
time
time
At low frequency
Amplitude
Frequency
E
Notes: 48

## Slide 49
Dynamic AFM mode – Harmonic Oscillator (phase)
Frequency
E
F
A
time
time
At resonance
Amplitude
Notes: 49

## Slide 50
Dynamic AFM mode – Harmonic Oscillator (phase)
Amplitude
Frequency
E
F
A
time
time
After resonance
Φ Phase
Notes: 50

## Slide 51
Cantilever
Sample
X/Y/Z Scanner
Optical
system
Cantilever amplitude
Magic box
Dynamic AFM mode (AM-AFM) – Operational scheme
Constant
amplitude
Notes: 51

## Slide 52
Dynamic AFM mode (AM-AFM) – in liquid, for biology
1) Part of the liquid participates (coupled) to cantilever motion
Notes: 52

## Slide 53
Dynamic AFM mode (AM-AFM) – in liquid, for biology
2) The resonance gets broader  decrease of Q factor. That’s because water has higher viscosity than air δliquid > δair
1) Part of the liquid participates (coupled) to cantilever motion
Notes: 53

## Slide 54
Dynamic AFM mode (AM-AFM) – in liquid, for biology
2) The resonance gets broader  decrease of Q factor. That’s because water has higher viscosity than air δliquid > δair
1) Part of the liquid participates (coupled) to cantilever motion
3) When excited mechanically,
the response of the tip is plenty of spurious peaks
Notes: 54

## Slide 55
Outline
1) Introduction to Scanning Probe Microscopy
2) The Force
3) Methods
Cantilevers
Static and Dynamic AFM
4) Instruments
5) Advanced and novel AFM methods
Notes: 55

## Slide 56
Cantilever deflection measurement
Optical beam deflection
G. Meyer and N. M. Amer, Appl. Phys. Lett. 53, 1045 (1988).
Notes: 56

## Slide 57
Optical beam deflection
G. Meyer and N. M. Amer, Appl. Phys. Lett. 53, 1045 (1988).
Cantilever deflection measurement
Notes: 57

## Slide 58
Optical beam deflection
G. Meyer and N. M. Amer, Appl. Phys. Lett. 53, 1045 (1988).
A
B
C
D
DEFL
DEFL
S
Cantilever deflection measurement
Notes: 58

## Slide 59
Optical beam deflection
G. Meyer and N. M. Amer, Appl. Phys. Lett. 53, 1045 (1988).
A
B
C
D
Lateral
L
Cantilever deflection measurement
DEFL
S
Notes: 59

## Slide 60
Interferometry
Optical beam deflection
G. Meyer and N. M. Amer, Appl. Phys. Lett. 53, 1045 (1988).
A
B
C
D
Lateral
D. Rugar, H. J. Mamin, R. Erlandsson, J. E. Stern, and B. D. Terris, Rev. Sci. Instrum. 59, 2337 (1988).
Interference of red and blue beams
Drawback:
DRIFT!!!
B. Hogenboom Imp. College London
Tip position measurement
L
DEFL
S
Notes: 60

## Slide 61
Calibrations
-1- Detector sensitivity
Acquire a force curve on a very rigid surface (i.e. mica or silicon)
Condition:
scanner extension (nm) = cantilever deflection (nm)
scanner extension (nm) = Φ x cantilever deflection (V)
Φ [nm/V]
Notes: 61

## Slide 62
Calibrations
-1- Detector sensitivity
Acquire a force curve on a very rigid surface (i.e. mica or silicon)
Condition:
scanner extension (nm) = cantilever deflection (nm)
scanner extension (nm) = Φ x cantilever deflection (V)
-2- Cantilever stiffness
Thermal method
Thermal energy = ½ KBT = Cantilever energy = ½ k deflection2
k [N/m]
Thermal energy
Φ [nm/V]
Notes: 62

## Slide 63
The « magic box »
Cantilever
Sample
X/Y/Z Scanner
Optical
system
Cantilever deflection
Magic box
Notes: 63

## Slide 64
The « magic box »
Cantilever
Sample
X/Y/Z Scanner
Optical
system
Cantilever deflection
Magic box
PID
=
Proportional – Integral – Derivative loop
output= P x (setpoint-input) +
I x ∫(setpoint-input)dt +
D x d(setpoint-input)/dt
OUTPUT
INPUT
SETPOINT
Defined
By AFM user
Notes: 64

## Slide 65
Tip excitation
-1- Piezo-dither mode
Present in all AFMs
Drawback:
Excitation in liquid
X. Xu et al., BioPhys. Journal,
95(5), 2520-2528, 2008
Notes: 65

## Slide 66
Tip excitation
-1- Piezo-dither mode
-2- Photothermal excitation
Provided by Asylum Research
Makes use of a 2nd laser (blue) to excite the cantilever base
Present in all AFMs
Drawback:
Excitation in liquid
X. Xu et al., BioPhys. Journal,
95(5), 2520-2528, 2008
From Asylum Research website
Notes: 66

## Slide 67
Lock-in Amplifier
Tip deflection
time
time
Lock-In amplifier
Input
Excitation
Piezo
Excitation
Amplitude
Phase
Outputs
Amplitude
Phase
Notes: 67

## Slide 68
Lock-in Amplifier
Lock-In amplifier
Amplitude
Phase
Outputs
Excitation
Input
=
Cantilever
deflection
High-pass filter
X
Low-pass filter
Function
generator
Notes: 68

## Slide 69
Outline
1) Introduction to Scanning Probe Microscopy
2) The Force
3) Methods
Cantilevers
Static and Dynamic AFM
4) Instruments
5) Advanced and novel AFM methods
Notes: 69

## Slide 70
High-Speed AFM
ELF3 LLPS droplets
Toshio Ando
Ando T. et al, PNAS 98(22), 12468-12472, 2001
Nessim Louafi
Walking myosin
Notes: 70

## Slide 71
High-Speed AFM
ELF3 LLPS droplets
Walking myosin
Toshio Ando
Ando T. et al, PNAS 98(22), 12468-12472, 2001
Nessim Louafi
FAST
XYZ
SCANNER
MEGAHERTZ
FAST (SMALL)
CANTILEVERS
Notes: 71

## Slide 72
Fast Force-Curves Imaging
Bruker: PeakForce Tapping
JPK: QI quantitative imaging
Asylum Research: Fast Force Mapping
From Bruker website
Height
Adhesion
Young modulus
elasticity
HERPES virus capside
from JPK QI website
Notes: 72

## Slide 73
Correlative AFM - Fluorescence
AFM + Fluorescence  Wide field illumination
AFM + Fluorescence  Confocal illumination
Notes: 73

## Slide 74
Correlative AFM - Fluorescence
AFM + Fluorescence  Wide field illumination
AFM + Fluorescence  Confocal illumination
TIRF
Fluorescence
AFM
Notes: 74

## Slide 75
Correlative AFM - Fluorescence
AFM + Fluorescence  Wide field illumination
AFM + Fluorescence  Confocal illumination
AFM
and
STORM
Dahmane, S. et al. , Nanoscale, 2019
Notes: 75

## Slide 76
Correlative AFM - Fluorescence
AFM + Fluorescence  Wide field illumination
AFM + Fluorescence  Confocal illumination
AFM
and
STORM
Dahmane, S. et al. , Nanoscale, 2019
Tip and excitation volume are aligned  the sample is scanned
Notes: 76

## Slide 77
Correlative AFM - Fluorescence
AFM + Fluorescence  Wide field illumination
AFM + Fluorescence Spectroscopy  Confocal illumination
AFM
and
STORM
Fernandes T. et al., Scientific Reports, 2020
Notes: 77

## Slide 78
TP Advanced Microscopy (AFM-confocal)
Mount an AFM cantilever inside an AFM, align a laser on the cantilever backside for the optical beam deflection detection
Notes: 78

## Slide 79
TP Advanced Microscopy (AFM-confocal)
Mount an AFM cantilever inside an AFM, align a laser on the cantilever backside for the optical beam deflection detection
Calibrate cantilever stiffness and optical beam deflection sensitivity
Notes: 79

## Slide 80
TP Advanced Microscopy (AFM-confocal)
Mount an AFM cantilever inside an AFM, align a laser on the cantilever backside for the optical beam deflection detection
Calibrate cantilever stiffness and optical beam deflection sensitivity
Approach the AFM tip in close vicinity of a biological membrane, align the AFM tip with a confocal microscope excitation spot
Notes: 80

## Slide 81
TP Advanced Microscopy (AFM-confocal)
Mount an AFM cantilever inside an AFM, align a laser on the cantilever backside for the optical beam deflection detection
Calibrate cantilever stiffness and optical beam deflection sensitivity
Approach the AFM tip in close vicinity of a biological membrane, align the AFM tip with a confocal microscope excitation spot
Acquire simultaneously AFM (QI mode) and confocal images of fluorescent biological membranes
Notes: 81

## Slide 82
TP Advanced Microscopy (AFM-confocal)
Mount an AFM cantilever inside an AFM, align a laser on the cantilever backside for the optical beam deflection detection
Calibrate cantilever stiffness and optical beam deflection sensitivity
Approach the AFM tip in close vicinity of a biological membrane, align the AFM tip with a confocal microscope excitation spot
Acquire simultaneously AFM (QI mode) and confocal images of fluorescent biological membranes
Treat AFM images using gwyddion software and confocal images using ImageJ
Hutin S. et al., PNAS, 2023
Notes: 82

## Slide 83
Tips for High-Resolution AFM imaging
1) The instrumental noise should be smaller than the aimed resolution.
2) The lateral resolution cannot be smaller than the size of the tip apex.
3) The decay length of the tip surface interaction should be smaller than the tip apex.
4) Imaging isolated three dimensional objects imposes limits on the lateral resolution.
5) The adhesion energy should be minimized. This is usually achieved by using sharp tips.
6) The applied force should be as small as possible.
7) Under identical conditions (tip radius and applied force), the stiffer the sample the better the lateral resolution.
Ricardo Garcia, Amplitude Modulation Atomic Force Microscopy, Wiley, 2010
Notes: 83

## Slide 84
Software for AFM image treatment
http://gwyddion.net/
OPEN SOURCE!
Notes: 84

## Slide 85
Some documents
Notes: 85

## Slide 86
Contacts
Luca Costa
luca.costa@cbs.cnrs.fr
Team Integrative Biophysics of Membranes (IBM)
https://integrativebiophysicsofmembranes.wordpress.com/
Notes: 86